PENN SCALE

CM-101     Price  Computing  Scale

Legal for Trade

Certificate #04-101A1

Specification

| Model | CM - 101 |
| --- | --- |
| Capacity | 30.00 x 0.01 lb / 15.000 x 0.005 kg |
| Display | LCD ( 5 / 5 / 6 ) |
| Op. Temp. | 0 C ~ 35 C / 32 F ~ 95 F |
| Dimension |  |
| Platter | 300 x 250 mm / 11.8 x 9.8 “ |
| Height | 110 mm / 4.3 “ |
| Length | 335 mm / 13.1 “ |
| Width | 320 mm / 12.5 “ |
| Power |  |
| DC | 4 V / 4 Ah Rechargeable Battery |
| AC | 120 V 50 / 60 Hz |
| Connecter | Standard AC power connecter |
| Charge Life | 120 hrs |

Calibration Procedure

( Authorized License Dealer only )

Step 1.   Turn the scale Power ON

Step 2.	  While the scale is in Testing Mode,

press TARE key

Step 3.	 Enter PASS-WORD

Step 4.	 When the scale displays

“ c a l , f 1 “, press TARE key

Step 5.	When the scale displays “ Uload”,

press TARE key with nothing on

the platter,

Step 6.	When the scale displays “ load “ ,

place 30 lb weight on the  Platter,

then press TARE

the scale will display “span pass”

< Note >

This procedure must be performed  ONLY  by the

authorized licensed personnel

If done by other, the field inspector from W & M

may NULLIFY  the validity of the scale

Method to check the voltage condition

Method 1:

After the self-checking when you switch on the scale, the price window will display: DC  *.** (the voltage now)

Method 2:

During the process of usage, you can press the * key and  00 key meanwhile, then the price window will display: DC  *.**  (the voltage now),then you can press any key to be back to the normal weighing condition.

Penn Scale Mfg. Co., Inc.

150 W. Berks Street

Philadelphia, PA 19122

Phone: (215) 739-9644

Fax: (215) 739-9640