CALIBRATION OF THE 7300+ SERIES BENCH SCALE

Quick Sheet – For complete explanation, See Manual!

Press ‘PRINT’ After Each Step to advance

NOTE:  The (3) buttons on the front panel have dual functions; when the scale is in the normal operating mode, they function exactly as they are labeled.  When in the CAL mode, they have the following duplicate purpose:

| Button Label: | Function during CAL Mode: | Description |
| --- | --- | --- |
| Zero | Zero/CLEAR Display | Clears setting back to ‘0’ |
| Units | Advance Display 0000 | Adds place holders <0 to advance |
| Print | ENTER | Choose and advances to next step |

Use the buttons to enter/change the following calibration data:

| Scale Display | Factory Preset Value | Explanation of CAL step |
| --- | --- | --- |
|  | To Enter CAL | Press and hold ZERO… displays ‘P X’ and ‘C X’ These document changes for the Audit Trail feature. |
| AC ? | Enter Access Code | 4-Digit code to restrict access to CAL functions. You have 10 seconds to enter correctly or CAL will abort. |
| CAL 10 | Exit Point* | Press ‘ENTer’ to continue or ‘0’ ‘ENTer’ to exit. |
| CAP 21 | Preset if factory matched | Enter the capacity in the Primary Units as a whole number. Press ‘ENTer’. |
| RES 22 | Primary Resolution | Entered at 1:10,000 for based on capacity; Accept by pressing ‘Print/ENTER’ or enter a new value then ‘ENTER’. |
| -0- 23 | Zero Range | Zero Range; enter as amount in Primary Units (factory preset is same as capacity. |
| A-0 13 | No | Yes if scale should re-zero on power up, No if not. |
| CNT 14 | Yes | CouNT; select No to disable count function. |
| CAL 20 | Exit Point* | Press ‘ENTer’ to continue or ‘0’ ‘ENTer’ to exit. |
| RES 22 | Primary Resolution | Entered at 1:10,000 for ‘Auto CAL feature; otherwise enter desired resolution for Primary weight units. |
| UnS 24 | 1 (Lbs) | Primary units; 1=Lbs or enter a new value from the chart |
| Prt 25 | Stable | Choose Stable/First stable/Unstable for PRINT condition |
| Cnd 26 | NO | Canadian specifications (choose Yes/No) |
| 0-t 27 | 0.25 of 1 grad | Zero Tracking value as % of one display graduation |
| CAL 30 | Exit Point | Press ‘PRINT’ to continue or change to ‘0’, ‘PRINT’ to exit. |
| 2 Un 31 | 2 Kgs or 3 G | Secondary units; 3= G (2-5 Lbs) 2= Kgs(10 lbs and up) |
| 2 rE 32 | 1:10,000 | Enter the Secondary display resolution (or based on CAP) or defaults to 1:10,000 displayed graduations. |
| CAL 40 | Exit Point | Press ‘PRINT’ to continue or change to ‘0’, ‘PRINT’ to exit. |
| FiL 41 | 3 | Display Filtering 0 (fastest) to 9 (slowest) pre-set at 3 |
| Nol 42 | Empty Platform | Empty Platform, press ‘PRINT’ |
| HLF 43 | ½ Capacity | First reference point, add weight, key amount, ‘PRINT’ Note: Prompts for 50% of CAP 21 in Primary Units. |
| FUL 44 | Full Weight | 2nd reference point, add weight, key amount, ‘PRINT’ Note: Prompts for 100% of CAP 21 in Primary Units. |
| Nol 42 | Empty Platform | Empty Platform, press ‘PRINT’ |
| CFG 50 | Exit Point | Press ‘PRINT’ to continue or change to ‘0’, ‘PRINT’ to exit. |

*AT EXIT POINTS, TO EXIT Choose CAL 0.  Press ‘PRINT’.

At Save? NO, use ‘SAMPLE SET’ to toggle to‘YES’

At Ent AC, enter new or old Access Code.  Press ‘PRINT’.

At Acc OK, unit will revert back to Weigh Mode.